
OphirLMMeasurement COM Object
21
Revision History


22
Overview


22.1
Introduction


32.2
USB Details of Devices


43
OphirLMMeasurement: Detailed Description


43.1
Registration


43.1.1
Additional Registrations


43.1.2
Troubleshooting Regsvr32


53.1.3
Versions of OphirLMMeasurement


63.2
Creating an Object


63.3
Threading issues


63.4
Error Codes


73.5
Methods and Events


73.5.1
Device Communications


73.5.2
General Information and Diagnostics


73.5.3
Sensor Configuration


73.5.4
Measurement Delivery


73.5.5
Trigger Settings


73.5.6
Events


73.5.7
Legacy Methods


83.5.8
Alphabetic Listing


103.6
Detailed Explanations of Methods


103.6.1
Conventions


103.6.2
Device Communications


133.6.3
General Information and Diagnostics


153.6.4
Sensor Configuration


273.6.5
Measurement Delivery


293.6.6
Trigger Settings


313.6.7
Events


323.6.8
Legacy Methods


334
Data Streams


334.1
Stream Formats


334.1.1
Standard


334.1.2
Turbo


334.1.3
Immediate


344.2
Pseudo code


344.2.1
Setup and Start


344.2.2
DataReady_Handler


354.2.3
Stop and Close Down


354.2.4
Configuring the Various Stream Modes


364.2.5
Configuring the Sensor’s Measurement Mode


364.2.6
GetData Status Codes





1 Revision History

	12
	05-Jan-21
	1. Support for X938 (partial)

2. Support for VS 2019

3. Added GetData Status Codes: Temperature, Alert Hot, Pulse Width, and PfP Energy

	11
	25-Mar-20
	4. Removed XP support

5. Added status to GetData method for reporting pulse frequency
6. Added information about the new interface ICoLMMeasurement2
7. Added functions for calibration date

8. Added functions for pulsed power pulse width

9. Added functions for low frequency power pulse frequency

	10
	11-Apr-19
	1. Added 844-PE-USB

2. Listed Python and Java as environments we supply a demo for

3. Added Creating an Object and Threading issues sections

	09
	22-Mar-17
	1. Added C++ demo information

	08
	30-May-16
	1. Added 841-PE-USB and 1919-R

2. Added status report to GetData method for reporting Filter status for sensors that self-detect their Filter State

3. Remove component for 32-bit applications on 64-bit Windows

	01
	30-Jul-14
	1. Initial Revision


2 Overview

2.1 Introduction

This document describes OphirLMMeasurement. This is the COM object provided by Newport that allows user applications to interface with the Newport 843-R-USB, 1919-R, 841-PE-USB, 844-PE-USB, and X938.
Besides their use as standalone, fully featured laser power/energy meters, Newport devices can also be used through OphirLMMeasurement. This COM object allows system integrators to integrate the measurement expertise of Newport equipment with in-house legacy analysis packages.
Demo client applications in VB.NET 2010/2019, VC# 2010/2019, VC++ 2010/2019, Python, Java, MatLab, and LabVIEW 8.6 are provided with the PMManager installation and can be found in C:\Program Files\Newport\PMManager 3.xx\Automation Examples\Com object.
In addition, OphirLMMeasurement has been tested with Delphi 5. In practice, OphirLMMeasurement can be used in any environment that allows interaction with COM automation servers (although it hasn’t been tested with tools other than those mentioned).

Special note for C++ users: Due the intricacy involved in using COM objects in C++, the C++ demo includes a wrapper class which provides a standard C++ interface to the COM object. To use this wrapper class, add OphirLMMeasurement.h and OphirLMMeasurement.cpp to your project, and #include "OphirLMMeasurement.h". The file program.cpp is a demo of how to use the class. See the comments in OphirLMMeasurement.h for more details.
Note: We provide a COM object and not a .NET automation because many clients’ applications do not use .NET and it is easy to use a COM object from .NET.
2.2 USB Details of Devices
The following lists the differences between the various Newport devices vis-à-vis their USB communications. OphirLMMeasurement hides all this from the client application but the system integrator must be aware of this to ensure he sets up the system with the appropriate hardware support.

	· 841-PE-USB
	One channel,
	Full Speed,
	Bus Powered,
	up to 5 interrupt IN endpoints

	· 844-PE-USB
	One channel,
	Full Speed,
	Bus Powered,
	up to 5 interrupt IN endpoints

	· 843-R-USB
	One channel,
	High Speed,
	Self powered,
	one interrupt IN endpoint

	· 1919-R
	One channel,
	High Speed,
	Self powered,
	one interrupt IN endpoint

	· X938
	Up to 2 channels,
	High Speed,
	Self-powered,
	one interrupt IN endpoint


3 OphirLMMeasurement: Detailed Description

This section describes registration of OphirLMMeasurement as well as its methods, events, and error codes.

3.1 Registration

PMManager PC software (3.10 and higher) registers OphirLMMeasurement as part of the installation process of the PMManager application. Registration is also possible via the Menu of the PMManager application. 

3.1.1 Additional Registrations
After installing the PMManager PC application on one PC, it’s possible to register the COM objects on additional PC’s without installing the full PMManager package. To do so, use the following steps

1. Copy the necessary files

· For 32-bit applications (running on 32-bit or 64-bit Windows): From C:\Program Files\Newport\PMManager 3.xx\Automation Examples\Com object\Redistributable\x86 to the target PC

· For 64-bit applications: From C:\Program Files\Newport\PMManager 3.xx\Automation Examples\Com object\Redistributable\x64 to the target PC. 

2. Run installophircom.bat file on the target PC. This will install the USB drivers as well as register the OphirLMMeasurement COM object and all other components necessary for full functionality.  The following is the contents of installophircom.bat
· rem  1) install COM (this is Ophir specific)

· regsvr32 OphirLMMeasurement.dll

· rem  2) install ophir devices drivers (Ophir specific INF files)

· swapinf StarlabDrivers 
3. If there is a need to remove the software from the target PC and to clean its registry, run RemoveOphirCOM.bat that is provided. The following is the contents of RemoveOphirCOM.bat
· rem  1) remove ophir devices drivers (Ophir specific INF files)

· swapinf StarlabDriversRemove 
· rem  2) remove COM (this is Ophir specific)

· regsvr32 /u OphirLMMeasurement.dll

4. The default COM object installation includes Ophir dialog windows informing the user to disconnect Ophir devices. If it is necessary to hide this message, then the install and uninstall batch files can be run silently. To do this, modify installophircom.bat and RemoveOphirCOM.bat as follows
· installophircom.bat
1. regsvr32 /s OphirLMMeasurement.dll

2. swapinf StarlabDrivers --silent
· RemoveOphirCOM.bat
1. swapinf StarlabDriversRemove --silent 
2. regsvr32 /u /s OphirLMMeasurement.dll

· Note: If you chose to install and remove silently, you must verify that the Ophir device has been disconnected before starting the process. Otherwise the installation will fail may leave unwanted information in the OC and you won’t know about it 
3.1.2 Troubleshooting Regsvr32

Microsoft reports that there are occasions where Regsvr32 doesn’t execute properly even though it is being run as an Administrator (see http://support.microsoft.com/kb/249873). The way to overcome this is to re-run the Regsvr32 command from an elevated command prompt.
To open an elevated command prompt, follow these steps:

· Windows 8.1 and Windows 8

Swipe in from the right edge of the screen, and then tap Search. Or, if you are using a mouse, point to the lower-right corner of the screen, and then click Search. Type Command Prompt in the Search box, right-click Command Prompt, and then click Run as administrator. If you are prompted for an administrator password or for a confirmation, type the password, or click Allow.
· Windows 7 and Windows Vista

Click Start, type Command Prompt or cmd in the Search box, right-click Command Prompt, and then click Run as administrator. If you are prompted for an administrator password or for a confirmation, type the password, or click Allow.
· Windows 10

Click the Start menu, type Command Prompt or cmd in the Search box, right-click Command Prompt, and then click Run as administrator. If you are prompted for an administrator password or for a confirmation, type the password, or click Allow.
3.1.3 Versions of OphirLMMeasurement 
The OphirLMMeasurement component works in Windows Vista, 7, 8, 8.1, and 10. We provide 2 variations.

· 32-bit component for 32-bit applications.

· 64-bit component for 64-bit applications.

3.2 Creating an Object
In order to use the library an object must be created. This object has methods which allow controlling the Newport measuring devices. To create the object, the ProgID OphirLMMeasurement.CoLMMeasurement or the GUID {B180613C-E514-4739-AADC-CAD4493910D7} can be used. How precisely to create a COM object depends on your development environment; see the provided samples for some possibilities.
Some environments require an interface name to be specified. For version 10.0 of the COM object (which comes with PMManager 3.50) the interface is ICoLMMeasurement2. The older interface ICoLMMeasurement will still work, but will not allow accessing the functions added in this version. Environments that do not require specifying the interface name will automatically use the newer interface.
3.3 Threading issues

The COM object provided is Apartment-threaded. This means (approximately) that it can be only used directly from the thread it was created in. The simplest way to use this library from multiple threads is to create a separate object in each thread. Since the state of the instruments is shared between objects, this has the same effect.

3.4 Error Codes

· 0x00000000 (S_OK): No Error

· 0x80004001 (E_NOTIMPL): Function Not Implemented

· 0x80070057 (E_INVALIDARG): Invalid Argument

· 0x80004005 (E_FAIL): Unspecified Failure

· 0x80040200: Device not opened
· 0x80040201: Device Already Opened

· 0x80040202: Drivers cannot be loaded

· 0x80040203: Load File Missing

· 0x80040300: Device Failed

· 0x80040301: Device Firmware is Incorrect

· 0x80040302: Sensor Failed

· 0x80040303: Sensor Firmware is Incorrect

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040306: This Sensor is Not Supported

· 0x80040307: Not Applicable in this Device

· 0x80040308: The Device is no longer Available

· 0x80040400: Save To Sensor Failed

· 0x80040401: Param Error

· 0x80040402: Failed to create Safe Array

· 0x80040403: Not Applicable in this Sensor

· 0x80040404: Value Out of Range

· 0x80040405: Command Failed

· 0x80040500: Stream Mode Not Started

· 0x80040501: A channel is in Stream Mode

The OphirLMMeasurement COM object follows the standard COM practice of returning an HRESULT from its methods. The HRESULT can be passed to GetErrorFromCode to get a descriptive string, or the standard COM GetErrorInfo function can be called (see documentation for your environment). Some client environments (such as VB6 and all .NET languages) do not return this HRESULT from the method call; instead they throw an exception when a method returns a failure HRESULT. This exception will normally contain within it the error number and the descriptive string.

3.5 Methods and Events
3.5.1 Device Communications

	Close
	CloseAll
	GetKnownWirelessDevices

	OpenUSBDevice
	OpenWirelessDevice
	ResetAllDevices

	ResetDevice
	ScanUSB
	ScanWireless


3.5.2 General Information and Diagnostics

	GetDeviceCalibrationDueDate
	GetDeviceInfo
	GetDriverVersion

	GetErrorFromCode
	GetSensorCalibrationDueDate
	GetSensorInfo

	GetVersion
	IsSensorExists
	


3.5.3 Sensor Configuration

	AddWavelength
	DeleteWavelength
	GetDiffuser

	GetFilter
	GetLowFreqPowerPulseFreq
	GetMeasurementMode

	GetPulsedPowerPulseWidth
	GetPulseLengths
	GetRanges

	GetThreshold
	GetWavelengths
	GetWavelengthsExtra

	ModifyWavelength
	SaveSettings
	SetDiffuser

	SetFilter
	SetLowFreqPowerPulseFreq
	SetMeasurementMode

	SetPulsedPowerPulseWidth
	SetPulseLength
	SetRange

	SetThreshold
	SetWavelength
	


3.5.4 Measurement Delivery

	ConfigureStreamMode
	GetData
	StartStream

	StopAllStreams
	StopStream
	


3.5.5 Trigger Settings

	GetExtTrigModes
	GetExtTrigOnOff
	GetExtTrigWindowTime

	SetExtTrigMode
	SetExtTrigOnOff
	SetExtTrigWindowTime


3.5.6 Events

	DataReady
	PlugAndPlay
	


3.5.7 Legacy Methods

	Read
	Write
	


3.5.8 Alphabetic Listing

	Method
	Communications
	Diagnostics
	Configuration
	Measurement
	Trigger
	Legacy

	AddWavelength
	
	
	(
	
	
	

	Close
	(
	
	
	
	
	

	CloseAll
	(
	
	
	
	
	

	ConfigureStreamMode
	
	
	
	(
	
	

	DataReady (event)
	
	
	
	
	
	

	DeleteWavelength
	
	
	(
	
	
	

	GetData
	
	
	
	(
	
	

	GetDeviceCalibrationDueDate
	
	(
	
	
	
	

	GetDeviceInfo
	
	(
	
	
	
	

	GetDiffuser
	
	
	(
	
	
	

	GetDriverVersion
	
	(
	
	
	
	

	GetErrorFromCode
	
	(
	
	
	
	

	GetExtTrigModes
	
	
	
	
	(
	

	GetExtTrigOnOff
	
	
	
	
	(
	

	GetExtTrigWindowTime
	
	
	
	
	(
	

	GetFilter
	
	
	(
	
	
	

	GetKnownWirelessDevices
	(
	
	
	
	
	

	GetLowFreqPowerPulseFreq
	
	
	(
	
	
	

	GetMeasurementMode
	
	
	(
	
	
	

	GetPulsedPowerPulseWidth
	
	
	(
	
	
	

	GetPulseLengths
	
	
	(
	
	
	

	GetRanges
	
	
	(
	
	
	

	GetSensorCalibrationDueDate
	
	(
	
	
	
	

	GetSensorInfo
	
	(
	
	
	
	

	GetThreshold
	
	
	(
	
	
	

	GetVersion
	
	(
	
	
	
	

	GetWavelengths
	
	
	(
	
	
	

	GetWavelengthsExtra
	
	
	(
	
	
	

	IsSensorExists
	
	(
	
	
	
	

	ModifyWavelength
	
	
	(
	
	
	

	OpenUSBDevice
	(
	
	
	
	
	

	OpenWirelessDevice
	(
	
	
	
	
	

	PlugAndPlay (event)
	
	
	
	
	
	

	Read
	
	
	
	
	
	(

	ResetAllDevices
	(
	
	
	
	
	

	ResetDevice
	(
	
	
	
	
	

	SaveSettings
	
	
	(
	
	
	

	ScanUSB
	(
	
	
	
	
	

	ScanWireless
	(
	
	
	
	
	

	SetDiffuser
	
	
	(
	
	
	

	SetExtTrigMode
	
	
	
	
	(
	

	SetExtTrigOnOff
	
	
	
	
	(
	

	SetExtTrigWindowTime
	
	
	
	
	(
	

	SetFilter
	
	
	(
	
	
	

	SetLowFreqPowerPulseFreq
	
	
	(
	
	
	

	SetMeasurementMode
	
	
	
	
	
	

	SetPulsedPowerPulseWidth
	
	
	(
	
	
	

	SetPulseLength
	
	
	(
	
	
	

	SetRange
	
	
	(
	
	
	

	SetThreshold
	
	
	(
	
	
	

	SetWavelength
	
	
	(
	
	
	

	StartStream
	
	
	
	(
	
	

	StopAllStreams
	
	
	
	(
	
	

	StopStream
	
	
	
	(
	
	

	Write
	
	
	
	
	
	(


3.6 Detailed Explanations of Methods
3.6.1 Conventions
3.6.1.1 Sensor Identification

The set of devices that OphirLMMeasurement supports includes multichannel devices which may have more than one sensor attached. In order to provide a common interface across all devices, OphirLMMeasurement uniquely identifies a sensor based on the device that it is attached to (hDevice) and the channel to which it is attached. For devices with only one channel <channel> should be 0. For a 2-channel device it can be 0 or 1, and for a 4-channel device 0, 1, 2, or 3.

3.6.1.2 Strings

In order to guarantee compatibility with the various Development Environments, OphirLMMeasurement returns strings in the BSTR type. Arrays of strings are returned as VARIANT *.
3.6.1.3 Common Parameters
	Parameter
	Description

	[in] LONG hDevice
	A handle to an open device, received from OpenUSBDevice.

	[in] LONG channel
	A channel number, 0 to 3, as applicable to the device.

	[out] VARIANT* options
	A list of values that a setting can take, in the form of an array of BSTR. If the setting is not applicable to the chosen sensor, the array will be empty.

	[out] LONG* index
	The 0-based index of the currently selected value of a setting. If the setting is not applicable to the chosen sensor, -1.

	[in] LONG index
	The 0-based index of a value of a setting. This must be in the range of the options valid for the setting, as returned in the <options> parameter of a Get… method.


3.6.2 Device Communications

These methods are responsible for managing the communications with Newport devices. This includes finding and registering the device(s) within OphirLMMeasurement, starting a communications session, and closing down communications with the device(s).
3.6.2.1 Close

	Name
	Close

	Parameters
	· [in] LONG hDevice

	Use
	This method ends the communication session with <hDevice>.

	Return Codes
	· 0x00000000 (S_OK): No Error
· 0x80040304: Bad Device Handle

· 0x80040501: A channel is in Stream Mode

	See Also:
	Device Communications, Alphabetic Listing, Common Parameters


3.6.2.2 CloseAll

	Name
	CloseAll

	Parameters
	None

	Use
	Equivalent to calling Close on all open devices.

	Return Codes
	· 0x00000000 (S_OK): No Error
· 0x80040501: A channel is in Stream Mode

	See Also:
	Device Communications, Alphabetic Listing


3.6.2.3 GetKnownWirelessDevices
	Name
	GetKnownWirelessDevices

	Parameters
	· [out] VARIANT* serialNumbers

	Use
	This method is an infrastructure stub for future support and at present only returns an error code

	Return Codes
	· 0x80004001 (E_NOTIMPL): Function Not Implemented

	See Also:
	Device Communications, Alphabetic Listing


3.6.2.4 OpenUSBDevice
	Name
	OpenUSBDevice

	Parameters
	· [in] BSTR serialNumber

· [out] LONG* hDevice

	Use
	Given <serialNumber> that was found with the ScanUSB method, will open a communication session with a device. <hDevice> is a handle which will be used for all communication with the device.

	Return Codes
	· 0x00000000 (S_OK): No Error
· 0x80070057 (E_INVALIDARG): Invalid Argument
· 0x80040201: Device Already Opened
· 0x80040300: Device Failed

· 0x80040301: Device Firmware is Incorrect

· 0x80040302: Sensor Failed

· 0x80040303: Sensor Firmware is Incorrect

· 0x80040306: This Sensor is Not Supported

· 0x80040308: The Device is no longer Available

· 0x80040405: Command Failed

	See Also:
	Device Communications, Alphabetic Listing


3.6.2.5 OpenWirelessDevice
	Name
	OpenWirelessDevice

	Parameters
	· [in] BSTR serialNumber

· [out] LONG* hDevice

	Use
	This method is an infrastructure stub for future support and at present only returns an error code

	Return Codes
	· 0x80004001 (E_NOTIMPL): Function Not Implemented

	See Also:
	Device Communications, Alphabetic Listing


3.6.2.6 ResetAllDevices
	Name
	ResetAllDevices

	Parameters
	None

	Use
	Closes and restarts all open devices. After calling this function ScanUsb should be called to redetect the devices.
Note: This function was only needed for older instruments. All current instruments detect automatically when the sensor is changed.
Note: This method will only function when no devices are streaming data (i.e. any channels which were put into Stream mode with the StartStream method have been taken out of stream mode with StopStream or StopAllStreams).

	Return Codes
	· 0x00000000 (S_OK): No Error
· 0x80040501: A channel is in Stream Mode

	See Also:
	Device Communications, Alphabetic Listing


3.6.2.7 ResetDevice
	Name
	ResetDevice

	Parameters
	· [in] LONG hDevice

	Use
	Restarts <hDevice>. Used after changing sensor attached to the device
Note: This function was only needed for older instruments. All current instruments detect automatically when the sensor is changed.
Note: This command method will only function when the device isn’t streaming data (i.e. any channels which were put into Stream mode with the StartStream method have been taken out of stream mode with StopStream or StopAllStreams).

	Return Codes
	· 0x00000000 (S_OK): No Error
· 0x80040200: Device not opened
· 0x80040304: Bad Device Handle

· 0x80040308: The Device is no longer Available

· 0x80040405: Command Failed

· 0x80040501: A channel is in Stream Mode

	See Also:
	Device Communications, Alphabetic Listing, Common Parameters


3.6.2.8 ScanUSB
	Name
	ScanUSB

	Parameters
	· [out] VARIANT* serialNumbers

	Use
	Scans the USB for all Newport devices that are presently attached to the PC.

<serialNumbers> is a pointer to a VARIANT containing an array of BSTR.

	Return Codes
	· 0x00000000 (S_OK): No Error
· 0x80040202: Drivers cannot be loaded
· 0x80040203: Load File Missing

· 0x80040402: Failed to create Safe Array

· 0x80040501: A channel is in Stream Mode

	See Also:
	Device Communications, Alphabetic Listing


3.6.2.9 ScanWireless
	Name
	ScanWireless

	Parameters
	· [out] VARIANT* serialNumbers

	Use
	This method is an infrastructure stub for future support and at present only returns an error code.

	Return Codes
	· 0x80004001 (E_NOTIMPL): Function Not Implemented

	See Also:
	Device Communications, Alphabetic Listing


3.6.3 General Information and Diagnostics

These methods provide general purpose information about the devices, sensors, and software that are presently attached and participating in the present communication session 

3.6.3.1 GetDeviceInfo

	Name
	GetDeviceInfo

	Parameters
	· [in] LONG hDevice

· [out] BSTR* deviceName

· [out] BSTR* romVersion

· [out] BSTR* serialNumber

	Use
	Gets name, ROM version and serial number of the device.

	Return Codes
	· 0x00000000 (S_OK): No Error
· 0x80040200: Device not opened
· 0x80040304: Bad Device Handle

· 0x80040308: The Device is no longer Available

	See Also:
	General Information and Diagnostics, Alphabetic Listing, Common Parameters


3.6.3.2 GetDeviceCalibrationDueDate
	Name
	GetDeviceCalibrationDueDate

	Parameters
	· [in] LONG hDevice

· [out] DATE* dueDate

	Use
	Get the date this instrument is due to be calibrated.

Note: Some instruments do not yet report the calibration due date and will return error 0x80040307.

	Return Codes
	· 0x00000000 (S_OK): No Error
· 0x80040200: Device not opened
· 0x80040304: Bad Device Handle
· 0x80040307: Not Applicable in this Device
· 0x80040308: The Device is no longer Available

	See Also:
	General Information and Diagnostics, Alphabetic Listing, Common Parameters, GetSensorCalibrationDueDate


3.6.3.3 GetDriverVersion

	Name
	GetDriverVersion

	Parameters
	· [out] BSTR* info

	Use
	Get version of drivers that OphirLMMeasurement uses.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80004005 (E_FAIL): Unspecified Failure

	See Also:
	General Information and Diagnostics, Alphabetic Listing


3.6.3.4 GetErrorFromCode

	Name
	GetErrorFromCode

	Parameters
	· [in] LONG errorCode

· [out] BSTR* info

	Use
	Given an error code, will pass back a string explanation of the error.

	Return Codes
	· 0x00000000 (S_OK): No Error

	See Also:
	General Information and Diagnostics, Alphabetic Listing


3.6.3.5 GetSensorCalibrationDueDate
	Name
	GetSensorCalibrationDueDate

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [out] DATE* dueDate

	Use
	Get the date this sensor is due to be calibrated.

Note: Some instruments do not yet report the calibration due date and will return error 0x80040307.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040307: Not Applicable in this Device
· 0x80040308: The Device is no longer Available

	See Also:
	General Information and Diagnostics, Alphabetic Listing, Common Parameters, GetDeviceCalibrationDueDate


3.6.3.6 GetSensorInfo
	Name
	GetSensorInfo

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [out] BSTR* serialNumber

· [out] BSTR* sensorType
· [out] BSTR* sensorName

	Use
	Gets serial number, type (thermopile, photodiode, pyroelectric), and name of sensor.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

	See Also:
	General Information and Diagnostics, Alphabetic Listing, Common Parameters


3.6.3.7 GetVersion

	Name
	GetVersion

	Parameters
	· [out] LONG* version

	Use
	Get version of the OphirLMMeasurement COM object.

	Return Codes
	· 0x00000000 (S_OK): No Error 

	See Also:
	General Information and Diagnostics, Alphabetic Listing


3.6.3.8 IsSensorExists

	Name
	IsSensorExists

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [out] VARIANT_BOOL* exists

	Use
	Check if a sensor is attached to <hDevice> <channel>.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040308: The Device is no longer Available 

	See Also:
	General Information and Diagnostics, Alphabetic Listing, Common Parameters


3.6.4 Sensor Configuration

These methods are used to configure the measurement parameters of the sensor. Some of these methods apply to all types of sensors (e.g. Range methods) and some are sensor-type specific (e.g. Filter methods apply to Photodiode sensors only, Diffuser methods are only for Pyroelectric and PD Energy sensors, etc.) 
3.6.4.1 AddWavelength
	Name
	AddWavelength

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [in] LONG wavelength

	Use
	Add <wavelength> to list of favorite wavelengths of the selected sensor.

Wavelength: Must be between the lower and upper wavelength limits as returned by the GetWavelengthsExtra method.
Note: This method is only applicable for sensors with a continuous spectrum. The set of lasers cannot be altered in sensors with a discrete set of lasers settings.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040403: Not Applicable in this Sensor

· 0x80040404: Wavelength Out of Range

· 0x80040405: Command Failed

· 0x80040501: A channel is in Stream Mode

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, DeleteWavelength, GetWavelengths, GetWavelengthsExtra, ModifyWavelength, SetWavelength


3.6.4.2 DeleteWavelength

	Name
	DeleteWavelength

	Parameters
	· [in] LONG hDevice

·  [in] LONG channel

·  [in] LONG index

	Use
	Delete wavelength of selected sensor specified in <index>.

Note: If the currently selected wavelength is deleted, a different wavelength is automatically chosen. If there is only one wavelength left it cannot be deleted.
Note: This method is only applicable for sensors with a continuous spectrum. The set of lasers cannot be altered in sensors with a discrete set of lasers settings

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040401: Param Error

· 0x80040403: Not Applicable in this Sensor

· 0x80040405: Command Failed

· 0x80040501: A channel is in Stream Mode

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, AddWavelength, GetWavelengths, GetWavelengthsExtra, ModifyWavelength, SetWavelength


3.6.4.3 GetDiffuser
	Name
	GetDiffuser

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [out] LONG* index

· [out] VARIANT* options

	Use
	Get the diffuser state of the selected sensor and what the diffuser options are.

For sensors that are not equipped with an adjustable diffuser, <index> will always return 0. This method applies to Pyroelectric and PD Energy sensors only.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040402: Failed to create Safe Array 

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, SetDiffuser


3.6.4.4 GetFilter

	Name
	GetFilter

	Parameters
	· in] LONG hDevice

· [in] LONG channel

· [out] LONG* index

· [out] VARIANT* options

	Use
	Get the filter state of the selected sensor and what the filter options are.

For sensors that are not equipped with an adjustable filter, <index> will always return 0. This method applies to Photodiode sensors only.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040402: Failed to create Safe Array 

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, SetFilter


3.6.4.5 GetLowFreqPowerPulseFreq
	Name
	GetLowFreqPowerPulseFreq

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [out] DOUBLE* value
· [out] DOUBLE* min
· [out] DOUBLE* max

	Use
	Get the current pulse frequency for low frequency power measurement mode, and the maximum and minimum supported values.

This method applies to Photodiode sensors only, and is only supported on some instruments.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available
· 0x80040403: Not Applicable in this Sensor

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, SetLowFreqPowerPulseFreq


3.6.4.6 GetMeasurementMode
	Name
	GetMeasurementMode

	Parameters
	· [in] LONG hDevice
· [in] LONG channel
· [out] LONG* index
· [out] VARIANT* options

	Use
	Get the measurement mode of the selected sensor and what the measurement modes are.

This method applies to all types of sensors.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040402: Failed to create Safe Array 

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, SetMeasurementMode


3.6.4.7 GetPulsedPowerPulseWidth
	Name
	GetPulsedPowerPulseWidth

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [out] LONG* value
· [out] LONG* min
· [out] LONG* max

	Use
	Get the current pulse width for pulsed power measurement mode, and the maximum and minimum supported values, in milliseconds.
This method applies to Thermopile sensors only, and is only supported on some instruments.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available
· 0x80040403: Not Applicable in this Sensor

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, SetPulsedPowerPulseWidth


3.6.4.8 GetPulseLengths

	Name
	GetPulseLengths

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [out] LONG* index
· [out] VARIANT* options

	Use
	Get the selected pulse length of the sensor in selected channel and what the pulse length options are.

This method applies to Pyroelectric and PD Energy sensors only.
The strings returned are the maximum pulse length (in time) for each of the various pulse length settings.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040402: Failed to create Safe Array 

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, SetPulseLength


3.6.4.9 GetRanges

	Name
	GetRanges

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [out] LONG* index
· [out] VARIANT* options

	Use
	Get the measurement range of the selected sensor and the list of all available measurement ranges.

This method applies to all types of sensors.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040402: Failed to create Safe Array 

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, SetRange


3.6.4.10 GetThreshold

	Name
	GetThreshold

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [out] LONG* index
· [out] VARIANT* options

	Use
	Get the selected threshold setting of the sensor in selected channel and what the threshold options are.

This method applies to Thermopile sensors measuring energy and to some Pyroelectric sensors as well.
For Thermopile sensors, the set of available thresholds will generally be “LOW MEDIUM HIGH”.
For Pyroelectric sensors with adjustable thresholds, there is a threshold setting per pulse length.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040402: Failed to create Safe Array 

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, SetThreshold


3.6.4.11 GetWavelengths

	Name
	GetWavelengths

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [out] LONG* index

· [out] VARIANT* options

	Use
	Gets the selected wavelength of selected sensor and list of all available wavelengths.
This method applies to all types of sensors. Response is the same format for sensors with discrete and continuous spectrums. For additional spectrum information use GetWavelengthsExtra.
For sensors with a continuous spectrum, these strings will be numeric, for sensors with discrete spectrums, these strings may contain non-numeric characters as well (e.g. VIS or NIR).

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040402: Failed to create Safe Array 

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, AddWavelength, DeleteWavelength, GetWavelengthsExtra, ModifyWavelength, SetWavelength


3.6.4.12 GetWavelengthsExtra

	Name
	GetWavelengthsExtra

	Parameters
	· [in] LONG hDevice
· [in] LONG channel
· [out] VARIANT_BOOL*modifiable
· [out] LONG* minWavelength
· [out] LONG* maxWavelength

	Use
	Get additional wavelength information of selected channel. 
If spectrum is continuous, then
· modifiable: True

· minWavelength, maxWavelength: minimum and maximum limits for wavelength adjustment with the  AddWavelength and ModifyWavelength methods
If spectrum is discrete then

· modifiable: False

· minWavelength, maxWavelength: Not applicable

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040403: Not Applicable in this Sensor

· 0x80040501: A channel is in Stream Mode

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, AddWavelength, DeleteWavelength, GetWavelengths, ModifyWavelength, SetWavelength


3.6.4.13 ModifyWavelength

	Name
	ModifyWavelength

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [in] LONG index

· [in] LONG wavelength

	Use
	Modify selected sensor’s favorite wavelength stored at <index> to the value <wavelength>.

Note: <wavelength> must be between the lower and upper wavelength limits as returned by the GetWavelengthsExtra method.
Note: This method is only applicable for sensors with a continuous spectrum. The set of lasers cannot be altered in sensors with a discrete set of lasers settings.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040401: Param Error

· 0x80040403: Not Applicable in this Sensor

· 0x80040404: Wavelength Out of Range

· 0x80040405: Command Failed

· 0x80040501: A channel is in Stream Mode

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, AddWavelength, DeleteWavelength, GetWavelengths, GetWavelengthsExtra, SetWavelength


3.6.4.14 SaveSettings
	Name
	SaveSettings

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

	Use
	Save all configuration parameters of selected sensor.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040400: Save To Sensor Failed

· 0x80040501: A channel is in Stream Mode

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters


3.6.4.15 SetDiffuser
	Name
	SetDiffuser

	Parameters
	· [in] LONG hDevice
· [in] LONG channel
· [in] LONG index

	Use
	Set diffuser state of selected sensor to <index>
This method applies to Pyroelectric and PD Energy sensors only.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040401: Param Error

· 0x80040403: Not Applicable in this Sensor

· 0x80040405: Command Failed

· 0x80040501: A channel is in Stream Mode

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, GetDiffuser


3.6.4.16 SetFilter

	Name
	SetFilter

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [in] LONG index

	Use
	Set filter state of selected sensor to <index>.
This method applies to Photodiode sensors only.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040401: Param Error

· 0x80040403: Not Applicable in this Sensor

· 0x80040405: Command Failed

· 0x80040501: A channel is in Stream Mode

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, GetFilter


3.6.4.17 SetLowFreqPowerPulseFreq
	Name
	SetLowFreqPowerPulseFreq

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [in] DOUBLE  value

	Use
	Set the pulse frequency for low frequency power measurement mode.

This method applies to Photodiode sensors only, and is only supported on some instruments.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040403: Not Applicable in this Sensor
· 0x80040404: Value Out of Range

· 0x80040405: Command Failed

· 0x80040501: A channel is in Stream Mode

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, GetLowFreqPowerPulseFreq


3.6.4.18 SetMeasurementMode

	Name
	SetMeasurementMode

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [in] LONG index

	Use
	Set measurement mode of selected sensor to <index>.
This method applies to all types of sensors.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040401: Param Error

· 0x80040403: Not Applicable in this Sensor

· 0x80040405: Command Failed

· 0x80040501: A channel is in Stream Mode

	See Also
	Sensor Configuration, Alphabetic Listing, Common Parameters, GetMeasurementMode


3.6.4.19 SetPulsedPowerPulseWidth
	Name
	SetPulseLength

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [in] LONG value

	Use
	Get the pulse width for pulsed power measurement mode, in milliseconds.

This method applies to Thermopile sensors only, and is only supported on some instruments.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040403: Not Applicable in this Sensor
· 0x80040404: Value Out of Range

· 0x80040405: Command Failed

· 0x80040501: A channel is in Stream Mode

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, GetPulsedPowerPulseWidth


3.6.4.20 SetPulseLength
	Name
	SetPulseLength

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [in] LONG index

	Use
	Set pulse length setting of the selected sensor to <index>. 

This method applies to Pyroelectric and PD Energy sensors only.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040401: Param Error

· 0x80040403: Not Applicable in this Sensor

· 0x80040405: Command Failed

· 0x80040501: A channel is in Stream Mode

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, GetPulseLengths


3.6.4.21 SetRange
	Name
	SetRange

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [in] LONG index

	Use
	Set measurement range of selected sensor to <index>.
This method applies to all types of sensors.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040401: Param Error

· 0x80040403: Not Applicable in this Sensor

· 0x80040405: Command Failed

· 0x80040501: A channel is in Stream Mode

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, GetRanges


3.6.4.22 SetThreshold

	Name
	SetThreshold

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [in] LONG index

	Use
	Set threshold of selected sensor to <index>.
This method applies to Thermopile sensors measuring energy and to some Pyroelectric sensors as well.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040401: Param Error

· 0x80040403: Not Applicable in this Sensor

· 0x80040405: Command Failed

· 0x80040501: A channel is in Stream Mode

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, GetThreshold


3.6.4.23 SetWavelength

	Name
	SetWavelength

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [in] LONG index

	Use
	Set wavelength of selected sensor to <index>.
This method applies to all types of sensors.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040401: Param Error

· 0x80040403: Not Applicable in this Sensor

· 0x80040405: Command Failed

· 0x80040501: A channel is in Stream Mode

	See Also:
	Sensor Configuration, Alphabetic Listing, Common Parameters, AddWavelength, DeleteWavelength, GetWavelengths, GetWavelengthsExtra, ModifyWavelength


3.6.5 Measurement Delivery
See section Data Streams for detailed explanations of the various streams of data that come out of the devices and through OphirLMMeasurement, setting up the stream, starting and stopping it, and parsing the response. 

3.6.5.1 ConfigureStreamMode
	Name
	ConfigureStreamMode

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [in] LONG mode

· [in] LONG nValue

	Use
	Configure measurement delivery for the selected device and channel. See section Data Streams for detailed explanations of setting up the stream, starting and stopping it, and parsing the response.
Value of mode

Meaning 

Acceptable values for nValue

0

Turbo mode

0: off; 1: on

1

Turbo frequency

Any frequency acceptable to the device

2

Immediate mode

0: off; 1: on



	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040403: Not Applicable in this Sensor

· 0x80040405: Command Failed

· 0x80040501: A channel is in Stream Mode

	See Also:
	Measurement Delivery, Alphabetic Listing, Common Parameters


3.6.5.2 GetData
	Name
	GetData

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [out] VARIANT* arrayValue

· [out] VARIANT* arrayTimestamp

· [out] VARIANT* arrayStatus

	Use
	This method uploads the data payload that the meter has delivered to OphirLMMeasurement. It should be called after the DataReady event has been received.
If called when no new data is available, then the arrays will be empty.
See section Data Streams for detailed explanations of setting up the stream, starting and stopping it, and parsing the response.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040402: Failed to create Safe Array

· 0x80040500: Stream Mode Not Started 

	See Also:
	Measurement Delivery, Alphabetic Listing, Common Parameters, Data Streams, GetData Status Codes


3.6.5.3 StartStream

	Name
	StartStream

	Parameters
	· [in] LONG hDevice
· [in] LONG channel

	Use
	Instruct the selected sensor to begin data streaming measurements to the PC. The format of the stream (buffered, turbo, immediate) can be configured with the ConfigureStreamMode method.
When StartStream is called, the COM object flushes the buffer of previously gathered measurements.

OphirLMMeasurement will fire the DataReady event when a new data packet has arrived and is ready for reading by the client application.
See section Data Streams for detailed explanations of setting up the stream, starting and stopping it, and parsing the response.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80004005 (E_FAIL): Unspecified Failure

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040403: Not Applicable in this Sensor

· 0x80040405: Command Failed

· 0x80040501: A channel is in Stream Mode

	See Also:
	Measurement Delivery, Alphabetic Listing, Common Parameters, Data Streams 


3.6.5.4 StopAllStreams

	Name
	StopAllStreams

	Parameters
	None

	Use
	Stops data streams of all sensors.
See section Data Streams for detailed explanations of setting up the stream, starting and stopping it, and parsing the response.

	Return Codes
	· 0x00000000 (S_OK): No Error 

	See Also:
	Measurement Delivery, Alphabetic Listing, Data Streams 


3.6.5.5 StopStream

	Name
	StopStream

	Parameters
	· [in] LONG hDevice
· [in] LONG channel

	Use
	Stops data streaming from selected sensor.
See section Data Streams for detailed explanations of setting up the stream, starting and stopping it, and parsing the response.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040500: Stream Mode Not Started 

	See Also:
	Measurement Delivery, Alphabetic Listing, Common Parameters, Data Streams 


3.6.6 Trigger Settings

These methods are not supported by Newport meters. 
3.6.6.1 GetExtTrigModes
	Name
	GetExtTrigModes

	Parameters
	· [in] LONG hDevice

· [out] LONG* index

· [out] VARIANT* options

	Use
	Gets the selected external trigger mode of the device and the set of available external trigger modes.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040307: Not Applicable in this Device

· 0x80040308: The Device is no longer Available

· 0x80040402: Failed to create Safe Array

	See Also:
	Trigger Settings, Alphabetic Listing, Common Parameters, External Trigger Modes


3.6.6.2 GetExtTrigOnOff
	Name
	GetExtTrigOnOff

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [out] LONG* index

· [out] VARIANT* options 

	Use
	To get the external trigger state.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040402: Failed to create Safe Array 

	See Also:
	Trigger Settings, Alphabetic Listing, Common Parameters


3.6.6.3 GetExtTrigWindowTime
	Name
	GetExtTrigWindowTime

	Parameters
	· [in] LONG hDevice
· [out] LONG* extTrigWindow

	Use
	Get the external trigger window time of <hDevice> in microseconds.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040307: Not Applicable in this Device

· 0x80040308: The Device is no longer Available

	See Also:
	Trigger Settings, Alphabetic Listing, Common Parameters, External Trigger Window


3.6.6.4 SetExtTrigMode

	Name
	SetExtTrigMode

	Parameters
	· [in] LONG hDevice
· [in] LONG index

	Use
	Set the external trigger mode of <hDevice>.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040307: Not Applicable in this Device

· 0x80040308: The Device is no longer Available
· 0x80040401: Param Error
· 0x80040405: Command Failed

	See Also:
	Trigger Settings, Alphabetic Listing, Common Parameters, External Trigger Modes


3.6.6.5 SetExtTrigOnOff

	Name
	SetExtTrigOnOff

	Parameters
	· [in] LONG hDevice

· [in] LONG channel

· [in] LONG* index

	Use
	Set the external trigger state.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040305: Bad Sensor Channel

· 0x80040308: The Device is no longer Available

· 0x80040401: Param Error

· 0x80040403: Not Applicable in this Sensor

· 0x80040405: Command Failed

· 0x80040501: A channel is in Stream Mode

	See Also:
	Trigger Settings, Alphabetic Listing, Common Parameters


3.6.6.6 SetExtTrigWindowTime

	Name
	SetExtTrigWindowTime

	Parameters
	· [in] LONG hDevice
· [in] LONG extTrigWindowTime

	Use
	Set the external trigger window time of <hDevice> in microseconds.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040307: Not Applicable in this Device

· 0x80040308: The Device is no longer Available
· 0x80040401: Param Error
· 0x80040405: Command Failed

	See Also:
	Trigger Settings, Alphabetic Listing, Common Parameters, External Trigger Window


3.6.7 Events
3.6.7.1 DataReady
	Name
	DataReady

	Parameters
	· [in] LONG hDevice
· [in] LONG channel

	Use
	This event is fired to inform the client application that the sensor at <hDevice> <channel> has delivered a new data payload. 
OphirLMMeasurement will not deliver the data to the client application until it calls the GetData method.
Note: Further DataReady events for the specified sensor will not be fired until OphirLMMeasurement’s flag has been cleared by the client application calling GetData.

	Return Codes
	N/A

	See Also:
	Measurement Delivery, Events, Alphabetic Listing, Common Parameters


3.6.7.2 PlugAndPlay
	Name
	PlugAndPlay

	Parameters
	None

	Use
	Fired when a device that has been found by the ScanUSB method has been disconnected from the system. The client application can rescan the USB to discover which device has been disconnected.

	Return Codes
	N/A

	See Also:
	Device Communications, General Information and Diagnostics, Events, Alphabetic Listing


3.6.8 Legacy Methods

At Newport, we highly recommend using OphirLMMeasurement for all measurement needs. However, end users that are working with RS232 on the 1919-R or X938 or just in general may want to work with the ASCII command paradigm of the instruments. Therefore these methods are included in OphirLMMeasurement.

The function sequence is always a call to Write followed by a call to Read, i.e. every Write must be followed by a Read, and every Read must be preceded by a Write. Failure to adhere to this sequence may cause communication trouble.

Note: For a full description of available commands, please see “Newport User Commands.pdf” included in the installation.
3.6.8.1 Read

	Name
	Read

	Parameters
	· [in] LONG hDevice
· [out] BSTR* reply

	Use
	Read <reply> from <hDevice>. <reply> is in the form of an ASCII string.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040308: The Device is no longer Available

· 0x80040405: Command Failed

	See Also:
	Legacy Methods, Alphabetic Listing, Common Parameters, Working with Legacy Methods


3.6.8.2 Write
	Name
	Write

	Parameters
	· [in] LONG hDevice
· [in] BSTR command

	Use
	Write <command> to  <hDevice>. <command> is in the form of an ASCII string.

	Return Codes
	· 0x00000000 (S_OK): No Error

· 0x80040200: Device not opened

· 0x80040304: Bad Device Handle

· 0x80040308: The Device is no longer Available

· 0x80040405: Command Failed

	See Also:
	Legacy Methods, Alphabetic Listing, Common Parameters, Working with Legacy Methods


4 Data Streams
See Measurement Delivery for the list of methods that manage the data delivery from the meter.
4.1 Stream Formats

OphirLMMeasurement provides three types of data stream formats from the various devices to the client application.
· Standard

· Turbo

· Immediate

4.1.1 Standard
This is the default data stream that OphirLMMeasurement uses. Data is buffered within OphirLMMeasurement and which then fires the DataReady event to inform the client application that new data is available. This event can be fired as often as once every 50 milliseconds depending on how quickly the data is coming in and how quickly the client application calls the GetData method to read out the data from OphirLMMeasurement.
Besides OphirLMMeasurement’s buffering, the Juno and Pulsar devices have hardware buffering mechanisms that enable them to reach very high measurement frequencies (10 KHz and 25 KHz respectively)

4.1.2 Turbo 

The USBI, Vega, and Nova-II devices reach 130Hz of measurement that they can deliver every pulse in Standard format. In order to reach higher frequencies (2000Hz) they have to work in Turbo Mode. Using the ConfigureStreamMode method, the client application can tell OphirLMMeasurement that it wants to work in Turbo Mode and what is the expected frequency of the laser being measured.
There is no need to use Turbo Mode in order to reach high frequency measurements with the Juno and Pulsar devices. 
4.1.3 Immediate

For clients that want to synchronize their device’s measurements with something else (for example energy of a pulse together with a beam profile of the pulse captured by a camera in another piece of software such as BeamGage), OphirLMMeasurement provides Immediate Mode streams. Using the ConfigureStreamMode method to put OphirLMMeasurement in Immediate Mode, whenever a new measurement arrives, OphirLMMeasurement immediately fires the DataReady event instead of buffering the reading. Do not use immediate mode with high frequency data as the data delivery mechanism will not be able to keep up.
Note: If ConfigureStreamMode is not called before StartStream, OphirLMMeasurement assumes that the client application is using Standard Mode. However if ConfigureStreamMode is called to setup OphirLMMeasurement for Turbo or Immediate Modes, it must be called again to return to Standard Mode.
4.2 Pseudo code

The following are pseudo code examples of how to gather data with the Continuous Send and Turbo Mode methods of data delivery. For a live example, see the VB and C# sample projects included in the PMManager installation directory.

The non-measurement methods are straightforward and can be understood directly from the sample code. Measurement methods involve working with events and need additional explanation

All Measurement Collection can be divided into three components

· Setup and Start

· Collection

· Stop and Close Down

4.2.1 Setup and Start
Start of Function

/* Scan the USB for 843-R-USB devices that have been attached to computer. 
Get array of devices serial numbers */
ScanUSB (snNum)



/* The following opens the first USB device and gets the device handle */

OpenUSBDevice (snNum(0), nHandle) 
/* Optional. Set the measurement parameters (i.e. Range, Wavelength, etc) if necessary */

Set Measurement Parameters () 

/* Optional – Configure measurement mode of sensor. The set of possible modes is dependent on the type of sensor in use. See Configuring the Sensor’s Measurement Mode */
SetMeasurementMode ()
/* Optional - Configure mode of stream. This section is different for the 3 different stream modes. See Configuring the Various Stream Modes  */

Configure the Stream Mode () 

/* Start the measurement process */ 
StartStream (nHandle, nChannel)

End of Function

4.2.2 DataReady_Handler
Start of Handler

/* This handler wakes up when the DataReady event is fired. The format of the data is identical in all 3 modes; the only difference being that dataArray.Length should always be 1 if set up for immediate mode

Measurements are delivered as three arrays containing the following data

· Timestamp(double) : Time of measurement (in milliseconds)

· value(double): Measurement (in watts or joules)

· status(long): See section GetData Status Codes for a detailed explanation of the status codes for the various measurement modes
GetData (nHandle, channel, dataArray, timeStampArray, statusArray)
For index = 0 to dataArray.Length - 1

LabelMeasurements(channel).Text =arrayValue(index)
 // value

LabelTime(channel).Text 
=
arrayTimestamp(index)
 // time in milliseconds
LabelStatus(channel).Text 
= 
arrayStatus(index)
 // status

Next
End of Hander

4.2.3 Stop and Close Down
This section is identical across all 3 stream modes

Start of Function

/* Stop gathering data */ 
StopStream (nHandle, nChannel) 
/* Close device*/
Close (nHandle)
End of Function

4.2.4 Configuring the Various Stream Modes
// values of mode and nValue parameters of the ConfigureStreamMode method

	Value of mode
	Meaning 
	Acceptable values for nValue

	0
	Turbo mode
	0: off; 1: on

	1
	Turbo frequency
	Any frequency acceptable to the device

	2
	Immediate mode
	0: off; 1: on


4.2.4.1 Standard

Start of Function

/* OphirLMMeasurement defaults to this state. Only necessary if a different type if stream mode was configured and now want to return to Standard */

ConfigureStreamMode (nHandle, nChannel, 0, 0) 
// Turbo Mode turned off
ConfigureStreamMode (nHandle, nChannel, 2, 0) 
// Immediate Mode turned off
End of Function


4.2.4.2 Turbo

Start of Function

/* Set expected frequency of laser and Turbo mode to on. Reminder: Turbo Mode is unnecessary for the Juno and Pulsar devices. Also, it is only applicable for Pyroelectric and PD Energy sensors. If Turbo frequency is set to greater than the max frequency that the sensor can handle, an error will be returned that contains the max possible Turbo frequency */

ConfigureStreamMode (nHandle, nChannel, 1, 5000) 
// Turbo frequency set to 5000Hz 
ConfigureStreamMode (nHandle, nChannel, 0, 1) 
// Turbo Mode turned on
End of Function


4.2.4.3 Immediate

Start of Function

/* Set Turbo Mode off and Immediate Mode on */

ConfigureStreamMode (nHandle, nChannel, 0, 0) 
// Turbo Mode turned off
ConfigureStreamMode (nHandle, nChannel, 2, 1) 
// Immediate Mode turned on
End of Function
4.2.5 Configuring the Sensor’s Measurement Mode

This section lists the various types of Newport sensors and which measurement modes they support. To get the appropriate list of modes for the sensor in use, call the GetMeasurementMode() method. To configure the sensor for measurement, call the SetMeasurementMode() method. Note that not all instruments support all modes.
	Sensor Type
	Measurement Modes

	Thermopile Sensor
	Power, Energy, Pulsed Power

	Thermopile Sensor with Beam Tracking
	Power, Energy, Pulsed Power, Power with Track

	Photodiode Sensor
	Power, Exposure, Low Freq Power

	Pyroelectric and PD Energy Sensors
	Power, Energy, Exposure


4.2.6 GetData Status Codes

The following is the list of status codes that can be returned with the measurements when calling the GetData() method. Other codes that may be reported are reserved for internal use or future expansion.
	Status Code
	Meaning
	When and Where

	0
	OK
	Power/Energy/Exposure measurements for all sensors

	1
	Overrange
	Power and Energy measurements for all sensors

	2
	Saturated
	Power and Energy measurements for all sensors

	3
	Missing
	Energy with Pyroelectric and PD Energy sensors on a Pulsar or X938 device

	4
	Reset state
	Single shot energy or Pulsed Power with a Thermopile sensor

	5
	Waiting
	Single shot energy or Pulsed Power with a Thermopile sensor

	6
	Summing
	Single shot energy or Pulsed Power with a Thermopile sensor

	7
	Timeout
	Single shot energy or Pulsed Power with a Thermopile sensor

	8
	Peak Over
	Single shot energy or Pulsed Power with a Thermopile sensor

	9
	Energy Over
	Single shot energy or Pulsed Power with a Thermopile sensor

	0x010000
	X measurement OK
	Track measurements with a BeamTrack sensor

	0x010001
	X measurement Error
	Track measurements with a BeamTrack sensor

	0x020000
	Y measurement OK
	Track measurements with a BeamTrack sensor

	0x020001
	Y measurement Error
	Track measurements with a BeamTrack sensor

	0x030000
	Size measurement OK
	Track measurements with a BeamTrack sensor

	0x030001
	Size measurement Error
	Track measurements with a BeamTrack sensor

	0x030002
	Size measurement Warning
	Track measurements with a BeamTrack sensor

	0x040001
	Filter State Change
	Photodiode sensors with built-in filter state detection reporting a change in filter state (in to out or vice-versa)

	0x050000
	Pulse frequency
	Power and Energy with Pyroelectric and PD Energy sensors

	0x100000
	Temperature
	Thermopile sensor

	0x200000
	Alert Hot
	Thermopile sensor

	0x300000
	Pulse Width
	Pulsed Power with a Thermopile sensor

	0x400000
	PfP Energy
	Pulsed Power with a Thermopile sensor


- 5 -


